SANE Magnet Review
Repaired polarized target magnet was successfully energized and put in persistent mode on 12/18/08 after the review, following the procedure described in the review report
On 12/19/08 the refrigerator was cooled down and CH2 was loaded in the insert

---
SANE Magnet Review
Thermal NMR signals from CH2 were observed on 12/19/08 at about 7:00 PM.
  - The frequency on 12/19 after ~24 h of field on was 213.4 MHz, B() = 5.012 T
  - The current at time of persistent mode switch was 77.300 A
  - Current was adjusted to 77.135 A to center the frequency at 213 MHz
  - the NMR frequency 212.98 MHz corresponds to 5.0019 T

---
Magnet Review Implementation
Modified operating procedures for energizing and de-energizing maget are posted on the SANE wiki under
  - http://hallcweb.jlab.org/experiments/sane/wiki/index.php/Magnet_Ramping

---
Magnet Review Implementation
Modified operating procedures for energizing and de-energizing maget are posted on the SANE wiki under
  - http://hallcweb.jlab.org/experiments/sane/wiki/index.php/Magnet_Ramping

---
Magnet Review Implementation
Modified operating procedures for energizing and de-energizing maget are posted on the SANE wiki under
  - http://hallcweb.jlab.org/experiments/sane/wiki/index.php/Magnet_Ramping

---
Magnet Review Implementation
Only 4 expert operators are allowed to modify or ramp the magnet current. These are: Don Crabb, Donald Day, James Maxwell, and Jonathan Mulholland
  - Names listed on procedure
The ramp rates are limited as follows:
  - 0 to 60 A 1.2 A/min
  - 60 to 72 A 0.6 A/min
  - 72 to 78.9 A 0.3 A/min Programmed on PS firmware & LabView
When in persistent mode the power supply is to remain on at the last setting -- operated in this mode in Dec. 2008 until quench; will continue in the future; step 19 of ramp up procedure
The helium level in the magnet is to be continually monitored.
  - saved in EPICS logger, displayed on target screen (PDP)
  - extensive target data logged by LabView converted to ASCII files for charting

---
Magnet Review Implementation
The EMF across the magnet terminals is to be monitored
  - installed a DVM that reads the coil voltage across taps a-k of “D” connector.
  - DVM displays the voltage locally and is readout by LabView via RS-232 for display on target screen, and logged in EPICS.
    - readout of all taps too complicated to implement on time
The CoO will be modified so that the magnet need only be off during work on the lower platform. It is to remain on during other access work. --
  - Modified COO reviewed by D. Skopik. Shift personnel are acknowledging reading it with signatures and e-mail. ACC staff have been informed and are aware of new procedure

---
Target Status
Magnet operated until quenched at 11:00 PM on 12/19/08
Target developed a leak, which prevented further operation
  - Leak required target warm up, but was not detected when target was disassembled
  - Cold cathode gauge reading chamber vacuum was off at that time, due to earlier rack power trip (gauge requires local reset):
    - no record of chamber vacuum during period before and sometime after quench. Gradual, unsuspected vacuum deterioration might have contributed to quench.
  - Damaged 4 K shield was found and rebuilt
  - Two failed rotary pumps and a damaged needle valve on the refrigerator have been replaced

---
Target Status (II)
Quench
  - Target insert encoder calibration was in progress at that time
  - Beam was being horizontally centered on the target, both rasters were on, envelope was 21 mm in diameter, current was 93 nA (ibcm1)
  - BPM data at time of quench and survey on 1/9/09 of refrigerator “nose” which holds insert indicate that
    - insert is 1.6 mm left of coil center, coil center is 2.6 mm left of beam
    - when beam was steered right the raster envelope was more than 1 mm away from the coil package at the moment the quench was already in progress (timed from magnet flow data)
    - combination of distance from coils, very short time in the vicinity of coils (< 1 s at  1.1 mm) and very low current density (0.3 nA/mm²) of rastered beam makes it unlikely beam may have caused quench
  - Future insert and beam centering will be done with reduced ( 10 mm) raster diameter.
  - SEM crate trip at time of quench may indicate power transient.

---
Target Status (III)
Target has been cooled and it may be ready for magnet powering today
All Dec. 2008 review measures have been implemented

---
Magnet and Beam Data
Magnet flow [slpm] (red; FI91115), beam current [nA] (green; ibcm1), and beam position [mm] (blue; IPM3H00B.XRAW) vs EPICS time (add 1 h for clock time) on 12/19/08.

---
Magnet and Beam Data
Beam position [mm] (IPM3H00A and IPM3H00B), Beam current [nA] (ibcm1) and Magnet flow [slpm] (FI91115) vs EPICS time (add 1 h for clock time) on 12/19/08.

---
Magnet and Beam Data
Magnet flow [slpm] (red; FI91115), beam current [nA] (green; ibcm1), beam position [mm] (blue; IPM3H00B.XRAW), LHe level (purple; LL91111) and BPM (cyan; IPM3H00A.XRAW) vs EPICS time (add 1 h for clock time) on 12/19/08

---
Magnet and Beam Data
Magnet flow [slpm] (red; FI91115), beam current [nA] (green; ibcm1), beam position [mm] (blue; IPM3H00B.XRAW), LHe level (purple; LL91111) and BPM (cyan; IPM3H00A.XRAW) vs EPICS time (add 1 h for clock time) on 12/19/08

---
SANE Magnet Review
Thermal NMR signals from CH2 were observed on 12/19/08 at about 7:00 PM.
  - The frequency on 12/19 after ~24 h of field on was 213.4 MHz, B() = 5.012 T
  - The current at time of persistent mode switch was 77.300 A, B(I) = 5.004 T
  - Current was adjusted to 77.135 A to center the frequency at 213 MHz
  - the NMR frequency 212.98 MHz corresponds to 5.0019 T

---
Only 4 expert operators are allowed to modify or ramp the magnet current. These are: Don Crabb, Donald Day, James Maxwell, and Jonathan Mulholland.
The ramp rates are limited as follows: 0 to 60 A 1.2 A/min
60 to 72 A 0.6 A/min
72 to 28.9 A 0.3 A/min
When in persistent mode the power supply is to remain on at the last setting.
The helium level in the magnet is to be continually monitored.
The EMF across the magnet terminals is to be monitored.
The CoO will be modified so that the magnet need only be off during work on the lower platform. It is to remain on during other access work.

---
Magnet Review Implementation
Modified operating procedures for energizing and de-energizing maget are posted on the SANE wiki under
  - http://hallcweb.jlab.org/experiments/sane/wiki/index.php/Magnet_Ramping

---
Magnet Review Implementation
Modified operating procedures for energizing and de-energizing maget are posted on the SANE wiki under
  - http://hallcweb.jlab.org/experiments/sane/wiki/index.php/Magnet_Ramping

---
Magnet Review Implementation
Modified operating procedures for energizing and de-energizing maget are posted on the SANE wiki under
  - http://hallcweb.jlab.org/experiments/sane/wiki/index.php/Magnet_Ramping