## Slide 1: UNH Target Lab Update
7/19/24

## Slide 2: Reliquifier Commissioning
Reliquifier commissioning was attempted in June
Unfortunately, a large portion of atmosphere got into the system which we’ve been debugging since
  We think this came from a combination of 1/3 of the cylinders not being properly clean on arrival + having the ballast open on the roots pumps

## Slide 3: Reliquifier Commissioning
Purity (%)
Cryo purifiers could not handle the ~75% purity helium in our system
  Spent ~2 weeks attempting to clean it.
  Minor success w/ small # of cylinders, but loss rate of ~90%
  QT recommended a membrane filter before the cryo filters if we want to clean <90% purity helium
16 Cylinders
1 Cylinder
70%
86%
78%
82.5%
100%
90%

## Slide 4: Reliquifier Commissioning
We have since emptied & cleaned 1/3 of our recapture cylinders
We successfully tested reliquification
  Mother dewar  transfer dewar  gas bag  liquid
  Recovery rate was ~91%
We successfully tested UHP gas helium recovery in our system w/ capture purity >99%
  Magnet boil-off
  Fridge boil-off
  Separator pumps
  Roots pumps
As of 7/18/24, we believe we have fixed the problem

## Slide 5: Reliquifier Commissioning

## Slide 6: Reliquifier Commissioning
We still need to either vent or clean the ‘junk’ helium in 2/3 of the storage tanks
Once complete, we are ready for physics running again
Expect to start physics running in ~1-2 weeks